Method for Merging Colored Dissolved Organic Matter (CDOM) Absorption Measurements from Long and Short Pathlength Spectrometers

Reference for CDOM merge method:

Trolley, G., Allen, K. and Mannino, A. (2026), An approach for merging colored dissolved organic matter absorption measurements from complementary benchtop spectrometers. Limnol Oceanogr Methods e70044. https://doi.org/10.1002/lom3.70044

Purpose:

Numerous instruments and configurations have been used to measure CDOM absorption (ag) samples collected in the field, with two of the common classes being short pathlength dual-beam benchtop spectrophotometers and long pathlength liquid waveguide capillary cell (LWCC) instruments. Short pathlength dual-beam spectrometers typically use sample cells to achieve pathlengths of 1–10 cm, which is optimal for samples and wavelengths where ag(λ) is high, but the system can struggle with low sensitivity for samples and wavelengths with low CDOM absorption, such as open ocean blue waters. Conversely, LWCC systems use capillary cells to achieve pathlengths of 200–500 cm, which provides high sensitivity for samples and wavelengths where ag(λ) is low, but can have issues with stability and baseline offsets, and can struggle at high ag(λ) due to enhanced relative internal stray light, which can lead to underestimations of the absorption coefficient in the far UV.

Method:

To make a merged CDOM product, which uses measurements from both types of instruments to deliver a consistent CDOM curve across a wider wavelength range, we use the value threshold approach described in Trolley et al 2026. The short pathlength instrument used is the Agilent double-beam Cary-100 Bio Ultraviolet–Visible scanning spectrophotometer, which uses sample cells with 10cm pathlength. The long pathlength instrument used is the WPI UltraPath UPUV waveguide system, which has a pathlength of ~2m. A given sample is processed separately with both instruments according to IOCCG protocols, and the calculated ag values are used to derive the merged product (Mannino et al., 2019).

The merged product uses Cary-100 data shorter than a certain wavelength (λ1) due to the high ag signal in the UV, uses Ultrapath data longer than a certain wavelength (λ2) due to the low ag signal in the red, and applies a weighted average of the two instruments in the transition zone between λ1 and λ2. The WPI Ultrapath spectrometer has a linear response up to 1.381 m-1 , meaning that the signal detected by the spectrometer ceases to linearly correlate with increased CDOM concentration at wavelengths where ag exceeds 1.381 m-1 (Cartisano et al., 2018). As such, λ1 is set as the first wavelength for a given sample where ag,ultrapath < 1.381 m-1 . For the Cary-100, previous work has measured the uncertainty of CDOM measurements on the order of 0.023–0.039 m-1 (Mannino et al., 2008). To achieve signal > 10x uncertainty using the conservative end of the reported range, λ2 is set as the last wavelength for a given sample where ag,Cary-100 > 0.4 m-1. Information on the two input files, the merge threshold values, and the merge threshold wavelength locations for a given merged sample are included in the comments of the SeaBASS file.

References:

Cartisano, C.M., Del Vecchio, R., Blough, N.V., 2018. A calibration/validation protocol for long/multi-pathlength capillary waveguide spectrometers. Limnol. Oceanogr. Methods 16, 773–786. https://doi.org/10.1002/lom3.10282

Mannino, A., Novak, M.G., Nelson, N.B., Belz, M., Berthon, F., Blough, N.V., Boss, E., Bricaud, A., Castillo, C.D., Vecchio, R.D., D’Sa, E.J., Matsuoka, A., Miller, R.L., Neeley, A., Tzortziou, M., Werdell, J., 2019. MEASUREMENT PROTOCOL OF ABSORPTION BY CHROMOPHORIC DISSOLVED ORGANIC MATTER (CDOM) AND OTHER DISSOLVED MATERIALS. IOCCG Ocean Opt. Biogeochem. Protoc. Satell. Ocean Colour Sens. Valid.

Mannino, A., Russ, M.E., Hooker, S.B., 2008. Algorithm development and validation for satellite-derived distributions of DOC and CDOM in the U.S. Middle Atlantic Bight. J. Geophys. Res. Oceans 113. https://doi.org/10.1029/2007JC004493